Heater controller command-response protocol
All command messages must be terminated with a two-byte CRC code followed by ETX {0x03}. All controller response messages are terminated with a two-byte CRC code followed by ETX. Except as a terminating character, and possibly as a byte in the CRC code, ETX may not appear in any message.

Command format details
Messages may contain multiple commands, separated by whitespace (spaces, tabs, new lines, or carriage returns).
[n] indicates that a whole number is expected.
[d] indicates that a decimal number is expected. A whole number is acceptable. Any digits provided beyond the expected precision are ignored. For example, "3", "-4.7", and "134.00" are all acceptable decimal inputs.

Response format details
In controller report templates, '#' represents a numeric digit, '.', a decimal point, '-', a negative symbol. Other template symbols represent non-numeric characters. When a reported number is positive, the negative symbol is replaced with a space, so the message length remains fixed.

Command	Description
r	Request a controller status report.

Response format template

"-###.# "... (repeated for a total of 16 channels = 112 bytes + "\r\n")
"M ###.## "... (repeated for a total of 6 heaters = 54 bytes + "\r\n")
"-###.# -###.# " (14 bytes + "\r\n")

The first line contains the 16 thermocouple channel temperatures, in °C, in channel order from 0-15.
The second line contains the operating mode M and power level for each of the 6 heaters, in channel order from 0-5. The mode is decoded as follows: '0' (zero) = off, 'm' = manual, 'a' = auto.
The third line contains the cold junction temperatures for the two multiplexers.

r[n]	[0..255] Start datalogging, or stop, if n is 0. Issue a controller status report every n seconds.

0	(Zero)	Turn the selected heater off.

a	Set the selected heater to auto mode (PID control), and turns it on if it’s off.

s[n]	Set the setpoint temperature, in °C, for the selected heater. [-200..1000]

m	Set the selected heater to manual mode, preserving its current power level, and turns it on if it’s off.
m[d]	[0..100.00] Set the selected heater's power level to the specified %. Switches the mode to manual, if the current mode is auto; but if the heater is off, it is not turned on.

h	Request the selected heater's status report.

Response format template (39 bytes + "\r\n")

"# -### M ###.## ###.## ## -##### -##### -##### -##### #####"

The values represented by the symbols in the template above are, in sequence:
Heater channel
Setpoint (the desired heater temperature, in °C)
Operating mode (M): '0' (zero) = off, 'm' = manual, 'a' = auto
Heater power level (percent)
Maximum power level (percent)
Thermocouple channel
Four PID parameters, typically tuned by step tests
PID gain Kc: stored as 100*Kc
PID integral coefficient Ci = 1 / Ti, stored as 100000*Ci
PID derivative coefficient Cd = Kc * Td = gain times dead time, stored as 10*Cd
PID preset gain Cpr = 1 / gp = reciprocal of process gain, stored as 1000*Cpr
Error (sum of active codes from the following list)
0	No error
1	ADC out of range
2	RS232 input buffer overflow
4	RS232 CRC error
8	Unrecognized command from RS232
16	Invalid device channel
32	Datalogging interval out of range (0..255)
64	Setpoint out of range
128	Control output power level out of range
256	Invalid thermocouple channel selected
512	Invalid thermocouple type
1024	Invalid configuration or configuration parameter
2048	ADC out of range
4096	PID commanded on heater with no thermocouple
8192	PV out of range (thermocouple disconnected?)

h[n]	[0..5] Select the specified heater and request its status report.

ht[n]	[0..15] Set the selected heater's thermocouple channel.

n[n]	[0..5] Select the specified heater.
t	Request the selected thermocouple's status report.

Response format template (17 bytes + "\r\n")

"## T ####.# #####"

The values represented by the symbols in the template above are, in sequence:
Thermocouple channel
Thermocouple type (T): '~' = none, 'K' = type K, 'T' = type T
Temperature, in °C
Error (see 'h' command for a list of values)

t[n]	[0..15] Select the specified thermocouple and request its status report.

tn[n]	[0..15] Select the specified thermocouple.

t~	Set the selected thermocouple's type to "none".

tK	Set the selected thermocouple's type to "K".

tT	Set the selected thermocouple's type to "T".

x[d]	[0..100.00] Set the selected heater's maximum % power level.

ck[n]	[0..32767] Set the selected heater's PID gain parameter (Kc). Specify the value in units of 0.01. For example, if Kc is 16.89, the command would be ck1689.

ci[n]	[0..32767] Set the selected heater's PID integral coefficient (Ci). Specify the value in units of 0.00001. For example, if Ci is 0.00243, the command would be ci243.

cd[n]	[0..32767] Set the selected heater's PID derivative coefficient (Cd). Specify the value in units of 0.1. For example, if Cd is 472.8, the command would be cd4728.

cr[n]	[0..32767] Set the selected heater's PID preset multiplier (Cpr). Specify the value in units of 0.001. For example, if Cpr is 0.652, the command would be cr652.

i	Toggle interference suppression. When enabled, temperature updates for a thermocouple are suppressed when its associated heater is receiving power.

z	Show program data.

Response format template

"Aeon Laboratories HC6 V.########-####" (37 bytes + "\r\n")
"S/N:####" (8 bytes + "\r\n")
"H:## T:###" (10 bytes + "\r\n")
"ADC:#####!" (10 bytes + "\r\n")

The first line contains the heater controller model number and the firmware version. The second line has the serial number. The third line shows the selected heater and thermocouple channels. The final line shows the ADC count for the currently selected thermocouple channel. The exclamation mark is present when interference suppression is enabled; otherwise, it is replaced with a space.